Set Up Instructions for Opus 5 and 10 Laser Fibers:

Following this procedure will result in your fiber not slipping in the handpiece. Make sure the fiber you have is like the one shown below.

Step 1: Strip 10 inches of the white, yellow or black outer jacket from the fiber. (Be careful not to nick the inner plastic covering).

`

Step 2: Strip 20mm or 3/4 inches (remove 7-10 mm each time until you remove 20 mm of the inner plastic jacket using your 200, 300 or 400 micron strippers. Removing 7-10 mm at a time assures minimum cladding damage. This will also assure that the fiber length is correct for your 90 degree Laser Glow Tips and will greatly minimize breaking .

Step 3: Cleave the fiber back about 1/4 inch and assure you have a clean break.

Step 4: With the cap in place slip the fiber thru the handpiece as shown until the stripped end is just at the front end of the handpiece.

Step 5: Put a 90 degree tip on the handpiece. Straighten the tip which will greatly increase your success at fiber feeding and minimize breakage thru until the protrusion you require is exposed usually 3-5mm for most procedures.

Step 6:  Rotate the cap until the fiber locks.

Step 7: Bend the tip to the required angle.

Step 8: When the handpiece gets close to the outer jacket simply remove the handpiece and slide the outer jacket off the inner fiber about 8-10 inches. Cut off the outer jacket being careful not to cut thru the inner plastic covering.

# Text boxes

Put tip here

Handpiece Internal View

Glass Core Strip 20 mm before cleaving

or 1.5 inches

Hard inner plastic covering the core

Outer Jacket remove 8 inches as handpiece gets close